## Slide 1: Floating Cable Trap Design Tool
Team: BoilerTrap
Folk Narongrit, Thejas Vishnu Ramesh, Joseph V. Rispoli Purdue University, West Lafayette, IN, USA
https://Github.com/MRBEL-Purdue/floating-cable-trap
Notes: Good afternoon everyone,
My name is Folk Narongrit from Purdue University, and our team will be presenting our open source floating cable trap design tool.

## Slide 2: What is a Floating Cable Trap?
Attenuate common-mode currents in the coaxial shields of MRI coils.
Floating cable traps attenuate signals in the coaxial shield without a physical connection
Normal Cable Traps
Floating Cable Traps
2
Notes:   First I want to give a quick overview of cable traps. They’re one of the most overlooked parts when designing the MRI coil system. The job of cable traps is to attenuate common-mode currents in the coaxial shields. Common-mode currents are not desired because it introduces noise and artifacts in the image and can increase the temperature of the cable.
  Normal cable traps are soldered onto the coil which has to be placed at the right location for optimal benefit.
  Floating cable traps don’t require any connections to the coaxial cable, allowing for flexible and optimal placement of the traps.
  On the top right image, You might have used a type of cable trap on USB cables – floating cable traps work in the same way.
  couple to nearby dielectrics (the body) a
      introduces resistive losses  Lowers SNR
  couples to other cables
      noise and artifacts in the reconstructed image
  increase in cable temperature
      Potential RF burns

## Slide 3: Design Equations
Open-Form Design Equations
Inductance
Capacitance
Resonance Frequency
(dielectric constant)
Closed-Form Design Equation
Normal Design
Process
This work
Notes: When we design cable traps, we normally use the three basic equations for a concentric cylinder to design the traps.
This way of designing cable traps does not guarantee optimal attenuation performance.
In fact, it does not even guarantee that the cable trap can be tuned to the Larmor frequency.
A cable trap designer might have to redo the calculations and refabricate the traps a couple of times to get it right.
In our work, we derived a single closed-form design equation by eliminating all intermediate variables from the problem.
We also included the tuning capacitance in the equation so the designer can predict which tuning capacitor to use.

## Slide 4: Demo – (Parametric Design Process)
Notes: Now I want to show a quick demo of the web application we wrote to compute and generate cable traps in a method we call the parametric design process.
You first input some constants, so we have the Larmor frequency, the known dielectric constant and magnetic permeability of the substrate,
and then the inner diameter, which should closely fit the cables you are using.
We then have some design constraints
Such as the capacitor values that you have in your lab, and the length of the cable trap, which is a physical constraint based on your application.
The program then shows you all the optimized design parameters, and an arrow will point at the most optimized cable trap parameters for the given length.
The program can then automatically generate a 3D model so that you can 3D print the cable traps.
[pause until video end]
On the top right are some example cable trap models that were produced through our program.

## Slide 5: Parametric Design Process
Code computes optimized trap’s dimensions that maximizes shield current attenuation
backup
Notes: [backup]
Now I want to show a quick demo of the web application we wrote to compute and generate cable traps in a method we call the parametric design process.
You first input some constants, so we have the Larmor frequency, the known dielectric constant and magnetic permeability,
and then the inner diameter, which should closely fit the cables you are using.
We then have some design constraints
Such as the capacitor values that you have in your lab, and the length of the cable trap, which is a physical constraint based on your experiments.
The program then shows you all the optimized design parameters, and an arrow will point at the most optimized cable trap for the given length.
The program can then automatically generate a 3D model so that you can 3D print the cable traps.
On the top right are some example cable trap models that was produced through our program.

## Slide 6: Optimized Design Combinations
Impossible Designs
3 T , K = 2.0, a = 0.95 cm
Isofrequency Contour Plots
9.4T, K = 2.0, a = 0.95 cm
Notes: Here we show some examples of the output plots at 3T on the left, and 9.4 T on the right.
You start seeing the benefits of using the program, where you can see certain designs that are not possible because they either cannot give you the desired frequency or the designs are physically impossible to manufacture.

## Slide 7: Performance
Bench Performance
MRI Performance
This work:
Empirically Designed:
3 T (128 MHz)
Phantom Only
(baseline)
Cable Only
Empirically Designed
This work
Notes: Here we show some of the experiments we have done comparing cable traps designed using our tool, versus one that was empirically designed using the standard equations.
The new cable traps was able to attenuate around 10 dB of common-mode current more than the old design.
While it does not sound like much, experiments at the scanner show that the new cable traps was able to recover SNR back to baseline values, and that signal homogeneity of the phantoms improved when using our cable traps [as opposed to one that was empirically designed.]

## Slide 8: Floating Cable Trap Design Tool
Download: https://github.com/MRBEL-Purdue/floating-cable-trap
Paper: Coming Soon!
Magnetic Resonance Biomedical Engineering Laboratory (MRBEL)
The Purdue Engineering MRI Facility
Team BoilerTrap
Folk Narongrit, Thejas Vishnu Ramesh, Joseph V. Rispoli Purdue University, West Lafayette, IN, USA
Thank you!
Notes: Our tool helps increase consistency and reduces fabrication time of cable traps while also improving SNR. There are many potential applications for MRI coils and other MRI technologies such as in interventional devices.
The next time you want to build a cable trap, feel free to go to our GitHub on QR code and try out our cable trap tool. We also have an upcoming paper outlining the design process so you can follow along in your lab.
Thank you very much