## Slide 1
Alexandre D’Astous1,2, Arnaud Bréhéret1, Eva Alonso-Ortiz1, Julien Cohen-Adad1,2,3,4
1. NeuroPoly Lab, Institute of Biomedical Engineering, Polytechnique Montreal, Montreal, QC, Canada
2. Centre de recherche du CHU Sainte-Justine, Université de Montréal, Montreal, QC, Canada
3. Functional Neuroimaging Unit, CRIUGM, Université de Montréal, Montreal, QC, Canada
4. Mila - Quebec AI Institute, Montreal, QC, Canada
Singapore, Canada
ISMRM abstract number: #3918
Notes: Hi all,
Thanks to the study group for having me
I will present ST, an open source software tool for B0 and B1 shimming

## Slide 2
Speaker Name: Alexandre D’Astous
I have no financial interests or relationships to disclose with regard to the subject matter of this presentation.
Declaration of Financial Interests or Relationships

## Slide 3: What is Shimming?
Making the field as close to the desired field as possible
Before shimming
After Shimming
Notes: What is shimming?
As you know, it’s making the field as close to the desired field as possible
For B0 for example its making it as close to the magnet strength as possible

## Slide 4
Loss of signal from T2* dephasing
Geometric distortions
Warped slice profiles
Disruption of frequency selective RF pulses
Ghosting, blurring and ringing artefacts
Widening of spectral linewidths in MRS
Magnet
Magnet
B0 field lines
Homogeneous
Inhomogeneous
B0 inhomogeneity effects
Notes: Why do we want to do that?
Because it creates loss of signal,…
Just to name a few

## Slide 5
Using shim coils
  Spherical harmonics
    Scanner shim coils
    Shim inserts
  Other coils
    Shim only multi-coils
    AC/DC, iPRES coils
Different shim volume
  Static (Volume)
  Dynamic (Slice-wise)
Different techniques
  Field mapping techniques
  Real-time shimming, maximum intensity
Custom AC/DC coil
Stockmann, J. P. et al.
How to shim?
Elster , A. D. MRIquestions.com
field map through time
Notes: How do we make it better,
We can use shim coils to create a field as close to the opposite of the inhomogeneities
We can do that using coils that have SH geometries (scanner which have klow order and /or shim insert which have higher orders)
There are also other coils such as shim only MC and ACDC or iPRE coils
We can also optimize different ROIs such as volume shim. But also dynamic shimming (slicewise)
Multiple techniques have also been proposed such as optimizing on fieldmaps, metods that relly on chosing the maximal intensity in magnitude images and real-time shimming which homogenizesthe field created by the breathing of patients
What’S the point, there are many techniques that can be used!

## Slide 6
Most vendors only support spherical harmonics and volume-wise optimization
  Their solutions are in “black boxes”
The software solutions require specific knowledge to implement
Each research group implements its own code
  Only works for a specific workflow
  Reinvents the wheel
  Hard to reproduce research
The problems
Notes: What are the problem?
Most vendors only support volume-wise SH shimming. It is often not enough
If you want to have access to these advanced shimmign solutions, you need to implement it
However, it requires
  Specific knowledge to implement
  And if the knowledge is available,
  Each research group implements their own version of the same thing that
    Often only works for specific sequences
    It reinvents the wheel and wastes time
    Produces hard to reproduce research

## Slide 7: Shimming Toolbox
1. McCarthy P. FSLeyes
2. D’Astous. A. et al., An open-source software toolbox for B0 and B1 shimming in Magnetic Resonance Imaging.
Notes: This is why we created ShimmingToolbox
It’s a software program that combines multiple tools to perform B0 and b1 shimming.
It can be installed on a computer that can be brought to the scanner room and compute shimming coefficients from a graphical user interface or from the command line using scripts to more easily reproduce experiments.
It includes tools to convert DICOM files to standardized BIDS dataset by leveraging dcm2niix and dcm2bids, it can create masks using BET and SCT or you favourite segmentation software and can compute B0 and B1shimming coefficients.
These tools allow to perform static, dynamic and real-time shimming using field map based methods, maximum intensity methods. Custom coils as well as scanner gradient and shim coils can be used to compute the shim coefficients allowing to shim using the scanner hardware as well as custom hardware.

## Slide 8: Main features
B0 field mapping / field unwrapping
B0 coil shim basis creation
B0 shimming
  Volume-wise, slice-wise, real-time
  Scanner coils (Spherical harmonics), custom coils
  Support for spherical harmonics for Siemens, Philips and GE
B1+ shimming
And more!
Notes: If you have other scanners we are looking forward to implementing more

## Slide 9: How we implemented it
Python
GitHub project
  Open-source
Website (https://shimming-toolbox.org/)
  Installation
  Tutorials
  Documentation
Available
  On the Command line
  As Graphical User Interface
Support for Mac, Linux and Windows
Notes: Shimming Toolbox is written in Python (you don’t even need matlab
Git project
  Open to the community
  Free to collaborate
We built a website
  Installation instructions
  Tutorials on how to use the software
  Documentation about the different APIs
ST is available on the command line for power users
But also available as a graphical user interface so that it can be used easily by everyone

## Slide 10: B0 slice-wise shimming results
Where: MGH Martinos Center for Biomedical Imaging (7T Siemens Terra)
Coil: 31-channel custom AC/DC coil
Structure: Brain
Sequence: EPI and field map
47% reduction of the STD
Notes: I’ll go fast around the result section in the interest of time but the goal is to show you that it works well
We collaborated with the MGH group where we did slice wise shimming in the brain
You can see on the left a 2nd order shim baseline done on the scanner compared to slice wise shimming using Shimming Toolbox
We see a great improvement and agreement between predicted and acquired shim images

## Slide 11: B0 dynamic shimming results
Notes: We also see less distortions in the EPI mages

## Slide 12: Slice-wise z-shimming
1. De Leener B., SCT: Spinal Cord Toolbox, an open-source software for processing spinal cord MRI data.
EPI acquisition using 21 different z-shim gradients, one for each volume
Mask is computed using the Spinal Cord Toolbox1
Signal intensity matrix in the spinal cord for each slice and each shim gradient
Notes: We also did slicewise zshimming in the spinal cord where we applied different z-shim gradients in EPI images and selected the one with highest signal intensity in the ROI

## Slide 13: Slice-wise z-shimming
Mean signal: 11%
EPI scan information
100 volumes
24 slices (5mm)
128 × 128 (1 × 1 mm2)
TE: 40 ms
TR: 2338 ms
Partial fourier: 7/8
Phase encode: AP
GRAPPA factor: 2
BW: 1220 Hz/Pixel
Signal in bottom slices: 80%
Notes: We saw great improvement in average signal intensity

## Slide 14: Slice-wise z-shimming
tSNR: 21%
Notes: Similar results in Tsnr measuremnts

## Slide 15: 𝑩𝟎 real-time dynamic xyz-shimming
Mean signal increased by 16%
STD decreased by 30%
Z
Where: Neuroimaging Functional Unit (3T Siemens Prisma)
Coil: Manufacturer’s gradient coil (1st order for real-time dynamic-xyz) and 2nd order (for global static)
Structure: Spinal cord
Sequence: Custom GRE
Notes: We alsodid realtime shimming which tried to improve the B0 field to correct for respiration induced B0 fluctuations and saw improvement in both the average signal and the standard deviation shown in red compared to standard 2nd order shim shown in black

## Slide 16: Static 𝑩𝟏+ shimming
Where: Montreal Neurological Institute (7T Siemens Terra)
Coil: 8Tx/20Rx channel custom coil
Structure: Spinal cord
Sequence: Pre-saturated TurboFLASH B1 mapping and MP2RAGE
UNI: 11% reduction in the CV
GRE: 40% reduction in the CV
UNI
GRE
*
Notes: B1 shimming also showed promising results
As you can see here, especially in the GRE images

## Slide 17: Conclusion
Open-source software for B0 and B1+ shimming
  Available for the command line and a Graphical User Interface
  Foster collaboration
  Make advanced shimming accessible and reproducible
Website: https://shimming-toolbox.org/
Notes: To conlude
ST is …
We have a couple of collaborations with research groups so if you are interested in trying it out or if you have any questiions. Let me know and I’ll be happy to help

## Slide 18: Acknowledgements
Abstract number: #3918
When: Wednesday 15h45
Session: Targeted Applications of Magnets, Shims & Gradient Design
Location: Exhibition hall
18
Thanks to all Shimming Toolbox
contributors
Notes: I would like to thank my collaborators as well as the Shimming Toolbox contributors who made this possible.